Manual for Microplastics Hyperspectral System

# Mechanical Setup
1. The z-achis light screw should be 15.5cm above the groundplate.
2. The y-achis light screw should be 25cm from the outer
3. Attach object table to xy-table by sliding it sideways into the groves.
4. Align the object tables right side with the blue line on the x-table. This way the object pit left side is aligned with the camera field of view, if the x-table is set to match the 3 red marks. The 3 green marks show the position for the right half of the object pit.
# Software Setup
1. Switch labscanner and camera on
2. Start “Lumo -Scanner”
## Tab Setup
1. Connect Sensor 1 – Camera
2. Sensor 1 – Set Capture Folder & Dataset prefix
3. Connect Motor 1
4. Find Limits Motor 1
5. Dataset naming: Set dataset naming format
## Tab Adjust
1. Find focusing target (QR-Code square) through moving target table with Motor – Arrows. Click “Sensor 1 position – Focus Grid – Set“ to store this position.
2. Focus camera until all edges are sharp
3. Find upper edge of white target and click “Sensor 1 position – White start – Set”
4. Move to the lower edge and click “Sensor 1 position – White stop – Set”. Make sure the view area stays at the same brightness.
5. With the camera view on the white target make sure the picture is not oversaturated (red areas in the view). In case of oversaturation reduce the exposure time in the “FX17 with dalsa” window. Don’t forget to click “Apply”
6. Locate the upper edge of the target area and click “Sensor 1 positions – Target start – Set”.
7. Move the x-table to center the framed box of the target pit in the camera view
8. Scanning speed calculation:
   select method: “Aspect ratio” and click run
   in the results popup select “scanner” (should be 0.9 – 1.0 mm/s)
## Tab Capture
Metadata
Set “Dataset name” – this will part of the file & folder name
Set “Description” – will be part of the metadatafile can be any text but no commas
# Measurement
1. Make sure box “Traffic lights” shows all green
2. Preview the capture area with button “Preview”
3. If preview is OK, click “Record”

# Auswertung
## Picking ROIs
1. Open Freelook, click “continue”.
2. Goto your data folder. Open “capture” folder and select the .raw file without DARKREF or WHITEREF prefix.
3. Double-Click “Band 1”
4. In the “Options” – menu click “Z profile”
5. Select the pixels you want to measure. Coordinates are shown on top off “Z Profile Plot”
6. Create a txt file in your data folder. Formatting has to be the same as the example-rois.txt